Table 2-6Data quality objectives and criteria for sediment chemistry measurements.

| QC Sample or Parameter | Frequency | Data Quality Objective/ Acceptance Criteria |
| --- | --- | --- |
| Polynuclear Aromatic Hydrocarbons |  |  |
| Initial Calibration | Prior to every batch sequence | 5-point curve. RRFb  30% RSDc |
| Continuing Calibration (Using mid-level calibration standard) | Every 10 field samples or 16 hours, whichever is more frequent | Δ RRFd  30% |
| Sediment SRM 1941 | One per batch | Values must be within ± 30% of true value on average for all analytes; not to exceed ± 35% of true value for individual analytes. |
| Procedural Blank | One per batch | No more than 2 analytes to exceed 5x target MDL unless analyte not detected in associated sample(s) or analyte concentration > 10x blank value. |
| Duplicate Sample | One per batch | RPDe  30% for 70% of all analytes |
| Matrix Spike | One per batch | %R = 60-125% |
| Surrogate (Internal) Standards | Every Sample | %R = 60-125% |
| Target MDLsa | -- | 1-5 ng/g (dry weight) |
| Polychlorinated Biphenyls and Pesticides |  |  |
| Initial Calibration | Prior to every batch sequence | 5-point curve. Standard curve correlation coefficient (r) > 0.9950 for all analytes |
| Continuing Calibration (Using mid-level calibration standard) | Every 10 field samples or 16 hours, whichever is more frequent | Δ RRFd  25% |
| Sediment SRM 1941 | One per batch | Values must be within ± 30% of true value on average for all analytes; not to exceed ± 35% of true value for individual analytes. |
| Procedural Blank | One per batch | No more than 2 analytes to exceed 5x target MDL unless analyte not detected in associated sample(s) or analyte concentration > 10x blank value. |
| Duplicate Sample | One per batch | RPDe  30% for 70% of all analytes |
| Matrix Spike | One per batch | %R = 60-125% |
| Surrogate (Internal) Standards | Every sample | %R = 60-125% |
| Target MDLsa | Sediment | 0.2 ng/g (dry weight) |
| Metals |  |  |

| Initial Calibration | Prior to every batch of samples | 5-point curve to assess linearity. RSDc for initial and subsequent calibrations must agree to within 2%. |
| --- | --- | --- |
| Continuing Calibration | Every 10 samples | See initial calibration |
| Procedural Blank | One per batch | Not to exceed 3x MDL unless analyte not detected in associated sample(s) or analyte concentration > 10x blank value. |
| Matrix Spike | One per batch | Method of additions used to assess signal suppression/enhancement |
| Sample Replicate Analysis | One triplicate analysis per batch | RSDc  25% |
| Sediment SRM 1646 | One per batch | Values must be within ±30% of true value on average for all analytes; not to exceed ±35% of true value for individual analytes. |
| Target MDLsa | -- | Concentrations in ppm (µg/g dry weight) by element: A1 (8 ppm); Fe (2.5 ppm); Mn (1.0 ppm); As (0.3 ppm); Cd (0.03 ppm by GFAA or 0.3 ppm by ICAP); CR (0.5 ppm); Cu (0.5 ppm); Pb (0.3 ppm); Hg (0.02 ppm); Ni (1.0 ppm); Se (0.3 ppm); Ag (0.02 ppm); Sn (0.8 ppm by GFAA or 8 ppm by ICAP); Zn (2.0 ppm); Sb (0.5 ppm by GFAA or 5.0 ppm by ICAP); Tl (0.5 ppm); Be (0.05 ppm); V (0.5 ppm). |

aMethod Detection Limits

bRelative Response Factor (RRF) = response of a given target analyte relative to the response of its associated internal standard.

cRelative Standard Deviation (RSD) = S/Mean x 100%

dChange in RRF (ΔRRF) = [(RRF of continuing calibration check - avg. RRF from initial calibration)  avg. RRF from initial calibration] x 100.

eRelative Percent Difference (RPD) = [Concentration 1 - Concentration 2]  Mean Concentration1,2 x 100.